## Slide 1: M501 Measurement of Hazardous Substances (including Risk Assessment) Day 3

## Slide 2: Learning Outcomes
Understand overnight questions
Understand types of sampling pumps & capture devices for dust, fume & fibre monitoring
Understand principles of workplace monitoring for dust, fumes & fibres

## Slide 3: Learning Outcomes (cont)
Understand equipment calibration & calculation of results
Review direct reading instrumentations and limitations of its use
Review common sources of error

## Slide 4
Dust, Fumes & Fibres

## Slide 5: Definition of Dust, Fumes & Fibres
Solid particles can exist as:
  Dust: solid material of varying sizes ( 0.1 – approx 100um)
  Fumes: produced when a solid is heated until a gas is generated and recondenses into solid or liquid particles ( typically < 1 um)
  Fibres: solid thread like filaments with a defined length to width ratio

## Slide 6: Typical Size Characteristics
Source: M Tranter 1999 –reproduced with permission

## Slide 7: Key Health Factors of Dust, Fumes & Fibres
Chemical composition of material
  Toxic effect : what is the toxicology of the material & the respective target organs?
Particle Size
  Where it deposits in the body : is it capable of penetrating to the alveoli or only the upper respiratory tract?

## Slide 8: Sampling Pumps
Many commercially available pumps
Most are small battery powered units which can be attached to a person
Operate at flow rates between 0.5 – 5.0 L/min however most particulate sampling is carried out at flow rates of 1.0 – 2.5 L/min

## Slide 9: Typical Sampling Pump
Source; University of Wollongong

## Slide 10: Useful Features of Pumps
Automatic flow control
Pulsation dampening
Capacity to operate at a reasonable backpressure
Reasonable flow range
Good battery capacity
Intrinsically safety

## Slide 11: Types of Operating Systems
Diaphragm
  Most common system currently in commercial use
  Requires pulsation dampening
Piston
  Not common but still used in some low flow pumps
  Pulsations are an issue
Rotary Vane
  Very precise engineering but quite heavy

## Slide 12: Schematic of a Diaphragm Sampling Pump
Source: BOHS – reproduced with permission

## Slide 13: Key Issues
Maintenance
  Must be performed regularly and recorded for each pump
  Check automatic flow compensation and internal inline filters
Battery charge
  Nickel-Cadmium batteries prone to “memory effect”. Cycling of pumps can overcome effect in most cases
  Use of appropriate chargers
Internal flowmeters
  Not accurate due to design flaw (one end must be open to atmosphere)

## Slide 14: Deposition Curves
In 1995 the International Standards Organisation (ISO) defined sampling conventions for particulates
  Inhalable fraction : inhaled through nose & mouth
  Thoracic fraction : penetrates beyond larynx
  Respirable fraction : penetrate to the alveoli

## Slide 15: ISO 7708-1995 Size Fractions for Particles
Source TSI Inc – reproduced with permission

## Slide 16: Cut Points of Size Fractions
Inhalable : typically all particles < 100um
Thoracic : all particles < 50um & 50% cut at 10um
Respirable : all particles < 16um & 50% cut at 4um

## Slide 17: Examples of Link to Health Effect
Coal dust:
  Disease: pneumoconiosis
  Effect: scaring of lung tissue therefore target organ is the lungs and thus respirable curve appropriate
Lead dust
  Disease: systemic poison
  Effect: toxic effect on blood system therefore target organ is the blood and thus inhalable curve appropriate

## Slide 18: Sampling Heads
Inhalable dust
  IOM sampling head (IOM)
  UKAEA 7 hole sampling head
  Conical inhalable sampler (CIS)
  SKC button sampler
  Pre-loaded cassettes

## Slide 19: IOM Sampling Head
Source: University of Wollongong

## Slide 20: The IOM Sampler Components
Cassette system
All collected dust is measured
Easily handled
No contact with filter
Multi fraction sampling with foam inserts
IOM sampler
IOM cassette
Transport clip
Front Cassette Filter Support O ring Body
cover front grid
Source :Airmet Scientific-reproduced with permission
Notes: <number>

## Slide 21: UKAEA 7 Hole Sampler
Source: HSE – reproduced with permission

## Slide 22: Conical Inhalable Sampler
Source: HSE – reproduced with permission

## Slide 23: SKC Button Sampler
Source: Airmet Scientific – reproduced with permission

## Slide 24: Pre-Loaded Cassette
Source: University of Wollongong

## Slide 25: Sampling Heads
Respirable dust
  BCIRA
  SIMPEDS
  Aluminium
  10mm Nylon (Dorr-Oliver)

## Slide 26: Operation of Miniature Cyclone
Source: HSE – reproduced with permission

## Slide 27: BCIRA (Higgins) Cyclone
Source: University of Wollongong
Notes: <number>
This was the original miniature cyclone and has only been included for historical interest. It was sometimes called the “Higgins” cyclone after one of its inventors.

## Slide 28: Respirable Dust Sampler (SIMPEDS)
Source: University of Wollongong

## Slide 29: Aluminium Cyclone
Source :Airmet Scientific – reproduced with permission

## Slide 30: Dorr-Oliver Cyclone
Source: University of Wollongong

## Slide 31: Sampling Heads
Thoracic Dust
  RespiCon sampler
  CIP 10 sampler

## Slide 32: RespiCon Sampler
Source; TSI Inc – reproduced with permission

## Slide 33: RespiCon Stage Impaction
Source; TSI Inc – reproduced with permission

## Slide 34: Special Sampling Heads
Asbestos & synthetic fibres
Diesel particulate
Rosin-based solder flux fume

## Slide 35: Asbestos & Synthetic Fibres
Source: University of wollongong

## Slide 36: Asbestos & Synthetic Fibres (UK)
Source: Gully Howard Technical – reproduced with permission

## Slide 37: Diesel Particulate Cassette
Source: Airmet Scientific – reproduced with permission

## Slide 38: Rosin-based Solder Fume Flux
Source: HSE – reproduced with permission

## Slide 39: Sample Train for Inhalable Dust
Pump
Connecting tube
Sampler
SAMPLER
CONNECTING
TUBE
PUMP
Source :Airmet Scientific – reproduced with permission
Notes: <number>

## Slide 40: Sampling Train for Respirable Dust
Source; University of Wollongong

## Slide 41: Position of Sampling Device
And MUST be taken in the Breathing Zone
300mm Hemisphere around the nose and mouth
Source :Airmet Scientific – reproduced with permission
Notes: <number>

## Slide 42: Sampling Train Connected to a Worker
Source :University of Wollongong
Remember:
Start pump & note start time-
Check flowrate during sampling-
At end of sample, stop pump &
note stop time-

## Slide 43: Key Points to Note
Need to ensure sampling tubing is secure
Need to collect appropriate information
Need to monitor sampling system several times during sampling period
Pre & post flow rates should be within +/- 5% as per “best practice”

## Slide 44: Type of Information to be Recorded
At commencement of sampling
  Sampler identification number
  Filter identification number
  Pump identification number
  Date & pump start time
  Initial flow rate of pump
  Workers name or description of static location

## Slide 45: Type of Information to be Recorded (Cont)
During sampling
  Description of task(s) undertaken during sampling period
  Risk control measures in place
  Atmospheric conditions
  Any other relevant data (e.g.-unplanned events)

## Slide 46: Type of Information to be Recorded (Cont)
At conclusion of sampling exercise
  Record the time
  Re-measure flow rate prior to switching off pump

## Slide 47: Calibration
Primary standards
  Traceable to a national standard
Secondary standards
  Requires calibration at regular intervals against a primary standard

## Slide 48: Primary & Secondary Standards
Primary standards
  Soap film meters
  Wet-test gas meters
  Bell spirometer
Secondary standards
  Electronic meters (some types considered primary standard in some countries)
  Rotameters
  Magnehelic gauges

## Slide 49: Soap Film Flow Meter
Source :Airmet Scientific – reproduced with permission

## Slide 50: Electronic Flow Meter
Source: University of Wollongong

## Slide 51: Rotameter
Source :Airmet Scientific – reproduced with permission

## Slide 52: Reading Rotameters
Source :Airmet Scientific – reproduced with permission

## Slide 53: Calibration System
Connect from pump to sampling head
Connect from sampling head to calibrator
Adjust flow to require flowrate
Source :Airmet Scientific – reproduced with permission

## Slide 54: Calibration with a Soap Film Flow Meter
Source: University of Wollongong

## Slide 55: Calibration with an Electronic Meter
Source: University of Wollongong

## Slide 56: Points to be Considered in Calibration
Use identical sampling head & filter to that used in field
Allow sample pump to stabilise
Measure flowrate of pump (3 consecutive readings within +/- 1 % of mean)
Take account of changes in environmental conditions such as altitude (if > 500m) & temperature (if >15°C)

## Slide 57: Suggested Calibration Schedules
All pumps: on use
Flow compensation
  Direct: 6 months if +/- 5% after 2 tests then 3 years
  Indirect Flow: 4 months if +/- 5% after 3 tests then 12 months

## Slide 58: Suggested Calibration Schedules
Rotameters
  Monthly for 3 months (+/- 3%) then 1 or 2 years depending on bore size
Soap Film Meter
  On commissioning
Electronic flow meters
  Monthly for 3 months (+/- 3%) then 6 monthly

## Slide 59: Suggested Calibration Schedules
Stopwatch
  6 Monthly
Balances
  1 point check monthly, 6 month repeatability check,
  36 months full range calibration by external authority

## Slide 60: Calculation of Results
To calculate the workers true exposure we require
  Total volume of air sampled
  Mass of contaminant on filter

## Slide 61: Total Volume Sampled
Volume (L) = Flowrate (L/min) X sample time (mins)
Volume (m3) = Volume (Litres)
1000
Note: 1 m3 = 1,000 L

## Slide 62: Mass on Filter
Mass (mg)= (post-pre weight of filter (mg))-blank (mg)

## Slide 63: Sample Concentration
Dust concentration (mg/m3) = Mass of contaminant (mg)*
    Sample volume (m3)
    * Corrected for blank

## Slide 64: Example Calculation
If flow rate was 2.2 L/min and sampling time was 7
hours and 42 minutes
Then:
Volume (L) = 2.2 x 462
=1,016.4
Volume (m3) = 1,016.4
1,000
= 1.0164 (1.02)

## Slide 65: Example Calculation
If the filter pre weight = 5.76 mg
and post weight = 7.84 mg
and blank = - 0.01mg
Corrected mass on filter = 7.84 - 5.75 - (- 0.01)
= 2.08 + 0.01
= 2.09 mg

## Slide 66: Example Calculation
Dust concentration (mg/m3) = 2.09
1.0164
= 2.056
= 2.1*
*rounded based on uncertainty of balance

## Slide 67: Direct Reading Instrumentation
Numerous instruments available on market
Most based on light scattering
Most optical based instruments over-respond in areas of high moisture
Emerging technologies addressing this problem

## Slide 68: Direct Reading Instrumentation
Very useful for
  Finding emission sources
  Measuring effectiveness of control technologies
  Highlighting dust issues to workers

## Slide 69: DustTrak
Source: TSI Inc – reproduced with permission
Very useful device in low moisture environments

## Slide 70: Personal Dust Monitor
Source; Thermo Fisher Scientific – reproduced with permission
Based on TEOM and may replace gravimetric sampling in some industries

## Slide 71: Tyndall Effect
Discovered by John Tyndall in mid 1800’s
  Principle can be used to highlight the presence of particles in the atmosphere
  Generally referred to as a “Dust Lamp”

## Slide 72: Dust Lamp
Source: BP International

## Slide 73: Most Common Sources of Error
The following, compiled by SKC Inc, are some common errors that have been found in dust sampling exercises
They are provided to highlight the need for care in all aspects of sampling

## Slide 74: Common Errors
Failure to clean cyclones before use
  To achieve the desired particle size separation, the internal parts of a cyclone must be clean
  Deposits of particulate matter adhering to the sides of the cyclone can alter the size-selection characteristics of the particulate penetrating the cyclone and collected on the filter

## Slide 75: Common Errors
Use of a pulsating pump for
collecting respirable dust samples
  Size-selective devices such as cyclones are affected by changes in flow rate.
  It is important to maintain a constant and non-pulsating flow rate to ensure correct size selection

## Slide 76: Common Errors
Use of area samples to assess
personal exposures
  The best estimation of a person’s exposure to a contaminant is obtained by placing the sampling equipment on the exposed individual. Area samples will be more difficult to defend as a reliable estimate of personal exposure

## Slide 77: Common Errors
Failure to use a constant flow pump
  Constant flow pumps automatically compensate for flow restrictions ensuring the flow rate is held constant
  Without this feature, users need to constantly monitor and manually adjust the flow rate to accurately measure air volume

## Slide 78: Common Errors
Failure to calibrate a pump properly
  Calibration, in air sampling, means to set and verify the flow rate
  Typically, this is done before and after every sample using a primary standard calibrator or using a secondary standard that has been calibrated to a primary standard

## Slide 79: Common Errors
Misuse of “self-calibrating” pumps
  Some sampling pumps available today have internal flow sensors that measure and display the flow rate
  These devices are secondary standards that should be verified with an external calibrator

## Slide 80: Common Errors
Failure to calibrate with recommended sampling media in-line
  Various types of sampling media (and the build up of dust) produce differing resistances to air flow (pressure drops) for which the pump must compensate.
  Standard methods require that pumps be calibrated within ±5% of the recommended flowrate with the sampling media in-line.

## Slide 81: Common Errors
Removing the grit pot from a cyclone
  The grit pot on a cyclone must be in place during calibration and sample collection
  An absent grit pot permits massive leakage and prevents proper collection of the desire dust size fraction

## Slide 82: Common Errors
Failure to sample at the design flow rate when using a cyclone sampler
  Each type of cyclone respirable dust sampler has a specific design flow rate that achieves the 50% cut-point
  Using a different flow rate will alter the collection efficiency curve including the 50% cut-point

## Slide 83: Common Errors
Inverting a cyclone during or after sampling
  The cyclone separator permits collection of smaller particles on the filter and removal of larger particles into the grit pot.
  Inversion of the cyclone causes larger particles to erroneously fall from the grit pot onto the filter material.

## Slide 84
Dust Practical

## Slide 85: Dust Practical - Overview
Learning outcomes
  Method selection
  Equipment selection
  Calibration
  Sampling
  Interpretation of data

## Slide 86: Dust Practical – Overview (cont)
Tasks
  Four (4) exercises
  Calculation of results
  Interpretation of data and report preparation
Group discussion

## Slide 87: Exercise 1 - Respirable Dust
Select appropriate equipment
Calibrate sampling train with electronic flow meter
Generate dust cloud with sander
Sample dust cloud
Recalibrate pump

## Slide 88: Exercise 2 - Inhalable Dust
Select appropriate equipment
Calibrate sampling train with soap film flow meter
Generate dust cloud with sander
Sample dust cloud
Recalibrate pump

## Slide 89: Exercise 3 - Evaluation of a Dust Cloud
Use MDHS 82 as a guide
Generate a dust cloud and observe “Tyndall effect”
Measure peak dust readings (if instrument available)
Discuss how this approach can be useful in developing a monitoring strategy

## Slide 90: Exercise 4 - Selection & Weighing of Filters
Inspect filters provided and select those suitable for monitoring respirable & inhalable dust
Equilibrate filters
Weigh filters

## Slide 91: Calculation & Interpretation of Data
Calculate workplace exposures from data provided
Establish level of risk within the workplace
Prepare an appropriate report

## Slide 92: Calculation & Interpretation of Data (cont)
Discuss aspects such as:
  monitoring strategy,
  any issues with data,
  outcome of assessment,
  limitations,
  possible recommendations
  any other relevant issues

## Slide 93: Review of Learning Outcomes
Understand overnight questions
Understand types of sampling pumps & capture devices for dust, fume & fibre monitoring
Understand principles of workplace monitoring for dust, fumes & fibres

## Slide 94: Review of Learning Outcomes (cont)
Understand equipment calibration & calculation of results
Review direct reading instrumentations and limitations of its use
Review common sources of error