## Slide 1: M504 – fibre counting

## Slide 2: Microscopy
Phase contrast microscope
Electron microscope

## Slide 3: Microscope type
Phase contrast microscope
At least 500 times magnification
  Objective lens 40x magnification
  Binocular eyepieces – at least 12.5x magnification
Stage micrometer
Walton-Beckett graticule in eyepiece
Test slide
Green filter

## Slide 4: Setting up microscope
Microscope must be set up correctly to achieve satisfactory performance
  A ‘Test slide’ (e.g. HSE Mk2 Test Slide) should be used to verify set up
  5 sets of lines must be visible

## Slide 5: Fibre counting
Filter preparation
  Filter must be ‘cleared’ (made transparent) and fixed to a microscope slide
  WHO method specifies use of acetone and triacetin to mount and clear the filters
  Method uses an ‘acetone vaporiser’

## Slide 6: Filter preparation
Filter is cleared with acetone vapour
Triacetin provides interface between collapsed filter and coverslip
Mounted slide is stable for many years
  Some accreditation bodies require slides to be kept for period of time e.g. UKAS require 6 months

## Slide 7: Acetone / triacetin filter clearing
Filter placed on slide (grid lines uppermost and parallel to edges of slide)
Slide and filter placed under outlet of heater unit
  Acetone injected into heater unit to produce stream of acetone vapour
  One drop triacetin added and coverslip placed gently onto surface
  Heat at about 50oC for 10-15 minutes to complete clearing

## Slide 8: Fibre counting rules
Phase contrast optical microscopy cannot positively identify individual fibres as asbestos
  Basis of WHO method is that all objects that conform to specified ‘size criteria’ are counted and assumed to be asbestos

## Slide 9: Countable fibres as defined by WHO 1997 method
Length: > 5µm
Width: < 3µm (average)
Length/width ratio greater than 3:1
If the fibre is attached to a particle it is assessed as if the particle does not exist and is counted if the visible part of the fibre meets the above definition

## Slide 10: Countable fibres as defined by WHO 1997 method
Fibres partially within the graticule
  Count ends of fibres
Split fibres
Fibre bundles
If more than 1/8th graticule area covered with dust / particles, area rejected for counting and another area chosen

## Slide 11
From: Determination of airborne fibre concentrations, WHO
(reproduced with permission)

## Slide 12
From: Determination of airborne fibre concentrations, WHO
(reproduced with permission)

## Slide 13
From: Determination of airborne fibre concentrations, WHO
(reproduced with permission)

## Slide 14
From: Determination of airborne fibre concentrations, WHO
(reproduced with permission)

## Slide 15
From: Determination of airborne fibre concentrations, WHO
(reproduced with permission)

## Slide 16
From: Determination of airborne fibre concentrations, WHO
(reproduced with permission)

## Slide 17
From: Determination of airborne fibre concentrations, WHO
(reproduced with permission)

## Slide 20: WHO method – Evaluation of air samples
Microscope set up for fibre counting
Initial examination of filter at low magnification to check uniformity of deposit and filter is undamaged
Filters examined at magnification of at least 500 times
Fine focus needs to be adjusted at each area examined
Graticule areas chosen at random over all of exposed filter area
  Area rejected if filter grid line obscures all or part of graticule area or if more than one eighth of graticule area is occupied by fibres and / or particles

## Slide 21: Number of Graticule Areas
Number of graticule areas examined will depend on sampling situation, volume sampled and quantification limit required
HSG248 (UK) gives the following guidance:
Personal sampling
  Generally 100 graticule areas are counted
Clearance, background, reassurance, leak monitoring
  200 graticule areas must be counted on 480 litre samples to achieve limit of quantification of 0.01 fibres/cm3
    If more than 480 litre is sampled, the number of graticule areas can be reduced proportionally

## Slide 22: Phase contrast optical microscopy method
Advantages
  Results available quickly – typically a filter can be cleared and analysed within about 30 minutes
  Equipment required is portable and can be taken to site to set up or in a small mobile laboratory
  Mounted filter is undamaged by analysis and is available for re-counting if required

## Slide 23: Phase contrast optical microscopy method
Disadvantages
  Method may not produce comparable results when used by different analysts
    Its precision is among the poorest of any occupational hygiene assessment method (WHO 1997)
  Differences may occur due to variations in
    Sampling methodology and sample preparation
    Equipment specification and set-up
    Largest potential source of differences is as a result of subjective effects associated with visual counting of fibres

## Slide 24: Phase contrast optical microscopy method
Limitations
  Method does not positively identify fibres as asbestos
    All structures conforming to specified size criteria are counted
  Upper width limit means some asbestos fibres are not counted
  Many fibres are too small to be visible by phase contrast optical microscopy
    Smallest fibres visible by phase contrast optical microscopy are about 0.2 micron width

## Slide 25: PCOM method - accuracy
Microscopists generally undercount dense deposits
In dusty environments other particles on the filter may interfere with accuracy of results and may make the filter unreadable
Systematic differences in fibre counts can occur between different microscopists within and between laboratories
  These intra- and inter-laboratory differences must be controlled by proper training and periodic quality checks

## Slide 26: Calculation of airborne fibre concentration
C = airborne concentration (fibres/ml)
N = number of fibres counted
D = diameter of exposed filter (mm)
V = volume of air sampled (litres)
n = number of graticule areas examined
d = diameter of graticule (µm)
Notes: <number>

## Slide 27: Fibre counting quality assurance
To ensure validity of results organisations that undertake asbestos air sampling and fibre counting should meet recognised standards
  Standards for competence of testing laboratories set out in ISO 17025
  Many countries have National accreditation schemes to assess the organisation against the ISO standard e.g.
    UKAS
    NATA
    AIHA

## Slide 28: Fibre counting quality assurance
All work must be carried out by competent trained staff using documented methods
Equipment suitable, regularly maintained and calibrated
Management systems in place to ensure accuracy, suitability and confidentiality of reports, together with secure storage and archiving of records and samples
Comprehensive auditing programme in place
System to investigate, correct and prevent incidents of non-conforming work and complaints

## Slide 29: Fibre counting quality assurance
An essential part of quality assurance is participation in internal quality control schemes and external proficiency testing (PT) schemes
  Particularly important for fibre counting due to large potential differences in results within and between laboratories
  Examples include RICE (UK), IHPAT (US), AFRICA (International)
  Slides sent to the testing laboratory for counting and results compared with the ‘reference’ value
  Satisfactory performance in PT schemes normally a pre-requirement for accreditation

## Slide 30: Fibre counting quality assurance
Example of proficiency testing scheme (RICE)
  ‘Good’ performance defined as 75% of counts within target band ‘A’ – for most slides this is a range of 0.65 to 1.55 as a ratio compared to the reference value
  Range may seem wide but reflects the subjective nature of fibre counting, uneven distribution on the filter and variations in microscope set-up.
    In fact the performance criteria have become more stringent over the years
  Proficiency testing schemes have improved consistency and comparability of fibre counting results both within and between countries by reducing systematic errors

## Slide 31: Electron microscopy
Phase contrast optical microscopy only gives a total fibre count
In some situations it may be required to establish whether counted fibres are asbestos or not e.g.
  Background testing when high fibre counts are found
  Leak testing where high ambient (non-asbestos) fibre levels are present
Electron microscopy is one technique that can be used to distinguish between different fibre types

## Slide 32: Electron microscopy
One problem is that acetone – triacetin mounted filters cannot be further analysed by electron microscopy
  Either filters must be cut in half before analysis or separate samples taken and reserved for electron microscopy
Electron microscopy uses an electron beam instead of a light beam
Used in conjunction with EDXA (Energy Dispersive X-ray Analysis) it enables identification of asbestos fibres by determining the chemical composition of the fibre

## Slide 33: Scanning electron microscopy (SEM)
Magnification is 2000x or greater
Fibre counting rules are the same as in the phase contrast optical microscopy method
  To achieve comparability with phase contrast optical microscopy method, fibres of less than 0.2 micron diameter are not counted
‘Countable’ fibres analysed by EDXA. X-rays generated within the fibre are detected and compared with standard spectra to classify the fibre as asbestos or non-asbestos
  This can be used to report a ‘discriminated’ fibre count as well as a total fibre count

## Slide 34: Transmission electron microscopy (TEM)
Can discriminate between fibre types of all sizes (as small as 0.01 micron in width)
Quantitative count of ‘fibres’ within randomly selected grid openings
As with scanning electron microscopy a discriminated fibre count can be derived
Scanning electron microscopy is used to give fibre counts comparable to phase contrast optical microscopy, whereas transmission electron microscopy is used more for research