## Slide 1
M501
Measurement of Hazardous Substances
(including Risk Assessment)
Day 4
Notes: <number>

## Slide 2: Today’s Learning Outcomes
Understand overnight questions
Understand the types of sampling techniques used for gas & vapour sampling
Understand the principles of workplace monitoring for gases & vapours including calibration of equipment and calculation of results
Review direct reading instrumentation & discuss limitations

## Slide 3
Sampling for Gases and Vapours

## Slide 4: Definitions
Gas- substance which is “air like’ but neither a solid or liquid at room temperature
Vapour-the gaseous form of a substance which is a solid or liquid at room temperature

## Slide 5: Types of Sampling
Grab or Instantaneous Samples
Source; BP International

## Slide 6: Types of Sampling
Short Term Samples
Source; BP International

## Slide 7: Types of Sampling
Long Term Samples
Source; BP International

## Slide 8: Types of Sampling
Continuous Monitoring
Source; BP International

## Slide 9: Sampling of Gases and Vapours
Whole of Air or Grab Sampling
Active sampling
  Absorption
  Adsorption
Diffusion or passive samplers
Direct reading instruments
Detector tubes

## Slide 10: Whole of Air or Grab Sampling
Collected
  Passively-evacuated prior to sampling
  Actively-by using a pump
Evacuated containers
  Canisters
  Gas bottles
  Syringes
Used when
  Concentration constant
  To measure peaks
  Short periods

## Slide 11: Whole of Air or Grab Sampling (cont)
Container preparation
  Cleaned
  Passivation eg Suma process
Compounds ideally
  Stable
  Recoveries dependent on humidity, chemical reactivity & inertness of container
  Down to ppb levels
  Landfill sampling

## Slide 12: Whole of Air or Grab Sampling (cont)
Gas bags e.g. Tedlar or other polymers
Filled in seconds or trickle filled
ppm levels
Source: Airmet Scientific – reproduced with permission

## Slide 13: Whole of Air or Grab Sampling (cont)
Sample loss issues:
  Permeation
  Adsorption onto bag
  Bag preparation
  Bag filling

## Slide 14: Whole of Air or Grab Sampling (cont)
Gas bags (cont)
Single use – cheap enough, but ??
If re use purge x 3 at least
Run blanks
Don’t overfill bag will take 3 times stated volume

## Slide 15: Active Sampling
Pump
Absorption
Adsorption – sorbent tubes eg
  Charcoal
  Silica gel
  Porous polymers – Tenax, Poropaks etc
  TD
Mixed phase sampling

## Slide 16: Active Sampling (cont)
Low volume pump –50 – 200 ml/min
Sample train
Calibration
Source: Airmet Scientific-reproduced with permission
Source: Airmet Scientific-reproduced with permission
Source: 3M Australia – reproduced with permission

## Slide 17: Active Sampling (cont)
Tube Holder
Source University of Wollongong

## Slide 18: Active Sampling (cont)
Break off both ends of a sorbent tube (2mm dia, or ½ dia of body)
Put tube in low flow adapter/tube holder
Make sure tube is in correct way around
Gas/Vapour Sampling Train
Source: Airmet Scientific – reproduced with permission

## Slide 19: Taking the Sample
Place sample train on person:
Start pump
Note start time
At end of sample:
Note stop time
Source :Airmet Scientific – reproduced with permission

## Slide 20: Active Sampling (cont)
Universal type pumps allow:
Up to 4 tubes at the same time – either running at different flow rates or with different tubes
Multi Tube sampling
To sample pump
3 way adaptor shown
Source :Airmet Scientific – reproduced with permission

## Slide 21: Absorption
Absorption – gas or vapour collected by passing it through a liquid where it is collected by dissolution in the liquid
Impingers
Source: University of Wollongong

## Slide 22: Absorption - Impinger Sampling Train
Source :Airmet Scientific – reproduced with permission

## Slide 23: Absorption (cont)
Collection efficiencies
  Size and number of bubbles
  Volume of liquid
  Sampling rate – typically up to 1 L/min
  Reaction rate
  Liquid carry over or liquid loss
  Connect in series
Need to keep samplers upright
Personal sampling awkward & difficult

## Slide 24: Absorption (cont)
Absorption derivatisation often used for:
  Formaldehyde collected in water or bisulphite
  Oxides of nitrogen – sulphanilic acid
  Ozone – potassium iodine
  Toluene diisocyanate – 1-(2- methoxy phenyl) piperazine in toluene

## Slide 25: Adsorption
Gas or vapour is collected by passing it over
and retained on the surface of the solid sorbent media
Main sorbent bed
Back up sorbent bed
Direction of sample flow
Source :Airmet Scientific – reproduced with permission

## Slide 26: Adsorption (cont)
Breakthrough:
Source :Airmet Scientific – reproduced with permission

## Slide 27: Adsorption (cont)
After sampling:
- remove tube
- cap the tube
- store, submit for analysis with details of sample
Don’t forget to send a blank with samples to laboratory
Source :Airmet Scientific – reproduced with permission

## Slide 28: Activated Charcoal
Extensive network of internal pores with very large surface area
Is non polar and preferentially absorbs organics rather than polar compounds
Typically CS2 for desorption

## Slide 29: Activated Charcoal (cont)
Limitations
Poor recovery for reactive compounds, polar compounds such as amines & phenols, aldehydes, low molecular weight alcohols & low boiling point compounds such as ammonia, ethylene and methylene chloride

## Slide 30: Silica Gel
Used for polar substances such as
Glutaraldehyde
Amines
Inorganics which are hard to desorb from charcoal
Disadvantage
Affinity for water
Desorption
Polar solvent such as water and methanol

## Slide 31: Porous Polymers & Other Adsorbents
Where gas & vapour not collected effectively with
charcoal or poor recoveries
Tenax – low level pesticides
XAD 2 – for pesticides
Chromosorb – pesticides
Porapaks – polar characteristics
Others:
Molecular sieves
Florisil for PCBs
Polyurethane foam for pesticides, PNAs

## Slide 32: Thermal Desorption
Superseding CS2 desorption especially in Europe
  Sensitivity
  Desorption efficiency
  Reproducibility
  Analytical performance

## Slide 33: Thermal Desorption (cont)
Thermal desorption tubes:
  ¼ inch OD x 3 ½ long stainless steel
  Pre packed with sorbent of choice
  SwageLok storage cap
  Diffusion cap
  Conditioning of tubes prior / after use
Sources: Markes International – reproduced with permission

## Slide 34: Thermal Desorption Unit with GC/MS
Sources: Markes International – reproduced with permission

## Slide 35: Collection Efficiencies of Adsorption Tubes
Temperature
  Adsorption reduced at higher temperatures
  Some compounds can migrate through bed
  Store cool box, fridge
  or freezer
Humidity
  Charcoal has great affinity for water vapour

## Slide 36: Collection Efficiencies (cont)
Sampling flow rate
  If too high insufficient residence time
Channeling
  If incorrectly packed
Overloading
  If concentrations / sampling times too long or other contaminants inc water vapour are present

## Slide 37: Mixed Phase Sampling
Solid, liquid, aerosol and gas and vapour phases.
  Benzene Soluble Fraction of the
  Total Particulate Matter
  for “Coke Oven Emissions”
  Impingers used for sampling
  of two pack isocyanate paints
  Aluminium industry – fluorides as particulate,
  or hydrofluoric acid as a mist or as gas.

## Slide 38: Treated Filters
Chemical impregnation including use for:
  Mercury
  Sulphur dioxide
  Isocyanates
  MOCA
  Fluorides
  Hydrazine

## Slide 39: Diffusion or Passive Sampling
Fick’s Law m = AD (c0 – c)
t L
where m = mass of adsorbate collected in grams
t = sampling time in seconds
A = cross sectional area of the diffusion path in square cm
D = diffusion coefficient for the adsorbate in air in square cm per second – available from manufacturer of the sampler for a given chemical
L = length of the diffusion path in cm (from porous membrane to sampler)
c = concentration of contaminant in ambient air in gram per cubic cm
c0 = concentration of contaminant just above the adsorbent surface in gram per cubic cm

## Slide 40: Diffusion or Passive Sampling (cont)
Source: HSE – reproduced with permission

## Slide 41: Diffusion or Passive Sampling (cont)
Every contaminant on every brand of monitor has its own
unique, fixed sampling rate
Source: 3M Australia – reproduced with permission

## Slide 42: Diffusion or Passive Sampling (cont)
Advantages
  Easy to use
  No pump, batteries or tubing & no calibration
  Light weight
  Less expensive
  TWA & STEL
  Accuracy ± 25% @ 95% confidence

## Slide 43: Diffusion or Passive Sampling (cont)
Limitations
  Need air movement 25 ft/min or 0.13m/sec
  Cannot be used for
    Low vapour pressure organics eg glutaraldehyde
    Reactive compounds such as phenols & amines
  Humidity
  “Sampling rate” needs to be supplied by manufacturer

## Slide 44: Diffusion or Passive Sampling (cont)
After sampling diffusion badges or tubes must be sealed and stored correctly prior to analysis
For example with the 3M Organic Vapour Monitors:
Single charcoal layer: Fig 1- remove white film & retaining ring. Fig 2 - Snap elution cap with plugs closed onto main body & store prior to analysis
Fig 1
Fig 2
Source: 3M Australia – reproduced with permission

## Slide 45: Diffusion or Passive Sampling (cont)
Those with the additional back up charcoal layer remove white film & snap on elution cap as above (Fig 3)
Separate top & bottom sections & snap bottom cup into base of primary section (Fig 4) and snap the second elution cap with plugs closed onto the back up section
Fig 3
Fig 4
Source: 3M Australia – reproduced with permission

## Slide 46: Diffusion or Passive Sampling (cont)
What can be typically sampled ?
Extensive range of organics
  Monitors with back up sections also available
Chemically impregnated sorbents allows
  Formaldehyde
  Ethylene oxide
  TDI
  Phosphine
  Phosgene
  Inorganic mercury
  Amines

## Slide 47: Calculation of Results
Active Sampling
Conc mg/m3 = mf + mr – mb x 1000
D x V
where mf is mass analyte in front section in mg
mr is mass analyte in rear or back up section in mg
mb is mass of analyte in blank in mg
D is the desorption efficiency
V is the volume in litres

## Slide 48: Calculation of results
Diffusion sampling:
Conc (mg/m3) = W (µg) x A
r x t
where W = contaminant weight (µg)
A calculation constant = 1000 / Sampling rate
r = recovery coefficient
t = sampling time in minutes
Conc (ppm) = W (µg) x B
r x t
where W = contaminant weight (µg)
B = calculation constant = 1000 x 24.45 / Sampling rate x mol wt
r = recovery coefficient
t = sampling time in minutes

## Slide 49: Direct Reading Instrumentation
Source: BP International

## Slide 50: Direct Reading Instruments
Many different instruments
Many different operating principles including:
  Electrochemical
  Photoionisation
  Flame ionisation
  Chemiluminescence
  Colorimetric
  Heat of combustion
  Gas chromatography
Many different gases & vapour
From relatively simple to complex

## Slide 51: Uses of Direct Reading Instruments
  Where immediate data is needed
  Personal exposure monitoring
  Help develop comprehensive evaluation programs
  Evaluate effectiveness of controls
  Emergency response
  Confined spaces

## Slide 52: Uses of Direct Reading Instruments (cont)
  For difficult to sample chemicals
  Multi sensors
  Multi alarms
  Stationary installations
  Fit testing of respirators
  Video monitoring

## Slide 53: Advantages
Direct reading
Continuous operation
Multi alarms
Multi sensors
TWA, STEL & Peaks
Data logging

## Slide 54: Limitations
Often costly to purchase
Need for frequent and regular calibration
Lack of specificity
Effect of interferences
Cross sensitivity
Need for intrinsically safe instruments in many places
Battery life
Sensors
  Finite life, poisoning, lack of range

## Slide 55: Cross Sensitivity of Sensors
Cross Sensitivity (CO Sensor)
H2S ~ 315
SO2 ~ 50
NO ~ 30
NO2 ~ -55
Cl2 ~ -30
H2 < 40
HCN 40
C2H4 90
Typical results from a challenge concentration of 100 ppm of each gas

## Slide 56: Filters for Contaminant Gases
H2S ~ 315 < 10
SO2 ~ 50 < 5
NO ~ 30 < 10
NO2 ~ -55 ~ -15
Cl2 ~ -30 < -5
H2 < 40 < 40
HCN 40 < 15
C2H4 90 < 50
Unfiltered
Filtered (typical)

## Slide 57: Other Limitations
Catalytic combustion detectors
  React with other flammable gases
  Poisoned by
    Silicones
    Phosphate esters
    Fluorocarbons

## Slide 58: Single Gas Monitor
Interchangeable sensors including:
O2, CO, H2S, H2, SO2, NO2, HCN
Cl2, ClO2, PH3
STEL, TWA, peak
Alarm
Data logging
Source: Industrial Scientific Inc – reproduced with permission

## Slide 59: Multigas Monitor
1 – 6 gases
Interchangeable sensors:
LEL, CH4, CO, H2S, O2, SO2,
Cl2, NO, ClO2, NH3, H2, HCl, PH3
STEL, TWA, peak
Alarm
Data logging

## Slide 60: Gas Badges
Two year maintenance free single
gas monitor
Sensors include CO, H2S, O2 and SO2
Turn them on & let them run out
Alarms
Some data logging ability
Source: Industrial Scientific Inc – reproduced with permission

## Slide 61: Photo Ionisation Detectors (PID)
Dependent on lamp ionisation potential
Typically non specific VOCs
or total hydrocarbons
  Some specific eg benzene, NH3, Cl2
Not for CH4 or ethane
Affected by humidity, dust,
other factors
Source: Airmet Scientific-reproduced with permission

## Slide 62: Flame Ionisation Monitor
Similar to, PID but flame
Non specific, broad range
Less sensitive to humidity &
other contaminants
Poor response to some gases
Needs hydrogen (hazard)
Source: Airmet Scientific-reproduced with permission

## Slide 63: Portable Gas Chromatograph
  Highly selective
  Range depends on type of detector used
  Complex instrument requiring
  extensive operator training
  Non continuous monitoring
Source: Airmet Scientific-reproduced with permission

## Slide 64: Infra-red Analyser
Organic vapours
Specific
Portable
Expensive

## Slide 65: Mercury Vapour Detectors
UV
  Interferences:
  Ozone
  Some organic solvents
Gold Film
  High cost
  Gold film needs regular cleaning

## Slide 66: Maintenance & Calibration
Source: Industrial Scientific Inc – reproduced with permission

## Slide 67: Guidelines for Using Gas Detection Equipment
Bump or challenge test
  Daily before use, known concentration of test gas to ensure sensors working correctly
Calibration
  Full instrument calibration, certified concentration of gas(es), regularly to ensure accuracy & documented
Maintenance
  Regular services provides reassurance instruments repaired professionally & calibrated & documented

## Slide 68: Typical Basic Instrument Checks
Physical appearance
Ensure instrument is within calibration period
Turn instrument on and check battery level
Zero the instrument
Bump test (functionality test) instrument
Clear the peaks

## Slide 69: Standard Gas Atmospheres
Primary Gas Standards
Are prepared from high purity 5.0 Gases (99.99999%) or 6.0 gases (99.999999%) by weighing them into a gas cylinder of known size
Secondary Gas Standards
Are prepared volumetrically from these using gas mixing pumps or mass flow controllers
Source: University of Wollongong

## Slide 70: Intrinsic Safety (cont)
IECEx Standards
Equipment for use in explosive or Ex areas eg
  Underground coal mines
  Oil refineries
  Petrol stations
  Chemical processing plants
  Gas pipelines
  Grain handling
  Sewerage treatment plants

## Slide 71: Intrinsic Safety (cont)
Gases, vapours, mists | Dusts | Explosive atmosphere is present
Zone 0 | Zone 20 | Most of the time
Zone 1 | Zone 21 | Some time
Zone 2 | Zone 22 | Seldom or short term
Classification of zones
Source: TestSafe – reproduced with permission

## Slide 72: Intrinsic Safety (cont)
Group 1 Equipment used underground
methane & coal dust
Group II Equipment used in other (above ground) hazardous areas
IIA - least readily ignited gases eg propane & benzene
IIB – more readily ignited gases eg ethylene & diethyl ether
IIC – most readily ignited gases eg hydrogen and acetylene
Gas or Explosive Groups

## Slide 73: Intrinsic Safety (cont)
Temperature classes
Group I Surfaces exposed to dust less than 150°C
Sealed against dust ingress less than 450°C
Group II
Temp Class | Max permissible surface temp °C
T1 | 450
T2 | 300
T3 | 200
T4 | 135
T5 | 100
T6 | 85
Source: TestSafe – reproduced with permission

## Slide 74: Intrinsic Safety (cont)
Levels of protection | Suitable for use in
“ia” | Zones 0, 20 (safe with up to 2 faults)
“ib” | Zones 1, 21 (safe with up to 1 fault)
“ic” | Zones 2, 22 ( safe under normal operation)
Levels of Protection & Zones
Source: TestSafe – reproduced with permission

## Slide 75: Intrinsic Safety Markings
Example Smith Electronics
Model TRE
Ex ia IIC T4
Cert 098X
Serial No. 8765
ia equipment suitable for zone 0 application
IIC equipment is suitable for Gas Groups IIA,IIB, IIC
T4 equipment is suitable for gases with auto ignition temp greater than 135°C

## Slide 76: Detector Tubes - Colorimetric Tubes
Change in colour of a specific reactant when in contact with a particular gas or vapour
Source: Dräger Safety – Reproduced with permission

## Slide 77: Advantages
Relatively inexpensive & cheap
Wide range of gases and vapours – approx 300
Immediate results
No expensive laboratory costs
Can be used for spot checks
No need for calibration
No need for power or charging

## Slide 78: Limitations
Interferences from other contaminants
Need to select correct tube & correct range
Results should NOT be compared to TWA
Correct storage
Limited shelf life

## Slide 79: Colour Tubes / Badges Available For
Instantaneous short term measurement
Long term measurements – pump
Long term measurements – diffusion
CHIP system
Based on colour reaction, but with digital readout of concentration

## Slide 80
Gas & Vapour Practical

## Slide 81: Gas and Vapour Practical - Overview
Learning outcomes
  Method selection
  Equipment selection
  Calibration
  Sampling
  Interpretation of data
Tasks
  Four (4) exercises
  Calculation of results
  Interpretation of data and report preparation
Group discussion

## Slide 82: Exercise 1 Sorbent Tube
Select appropriate equipment
Calibrate sampling train with electronic flow meter
Release / generate organic vapour
Sample “test” atmosphere
Recalibrate pump

## Slide 83: Exercise 2 Direct Reading Instrumentation
Select appropriate equipment
Establish limitations of instrument
Establish calibration requirements
Sample “test” atmosphere

## Slide 84: Exercise 3 Colorimetric Tubes
Select appropriate tube(s) and sampling pump measurement of organic vapours
Check operation of sampling pump
Sample “test” atmosphere
Take concentration readings
Notes: <number>

## Slide 85: Exercise 4 Diffusion OVM Badge
Select appropriate diffusion badge for organic vapours
Prepare badge for sampling
Sample “test” atmosphere
Conclude sampling and store collection device

## Slide 86: Calculation & Interpretation of Data
Calculate workplace exposures from data provided
Establish level of risk within the workplace
Prepare a short report. Discuss aspects such as:
  monitoring strategy,
  any issues with data,
  outcome of assessment,
  limitations,
  possible recommendations
  any other relevant issues

## Slide 87: Review of Today’s Learning Outcomes
Understand overnight questions
Understand the types of sampling techniques used for gas & vapour sampling
Understand the principles of workplace monitoring for gases & vapours including calibration of equipment and calculation of results
Review direct reading instrumentation & discuss limitations